## Calibration Certificate
CALIBRATION CERTIFICATE
Certificate No.: |  |  | CY/2607/ | 219.03 |  |  |  | Calibration Date: |  |  | 2026-07-31
ULR No.: |  |  | CC324926000000567F |  |  |  |  | Issue Date: |  |  | 2026-08-01
1. Calibrated For (Name & Address): | [7 empty] | Job No.: |  |  | CY/ | 219.03
| International Tractors Limited, | [6 empty] | Item Received Date: |  |  | 2026-07-29
| Village Chak Gujran, P.O. Piplawala, Jalandhar Road, | [6 empty] | Recommended Date
| Hoshiarpur, Punjab-146022 | [6 empty] | for Next Calibration: |  |  | 2027-07-30
2. Calibration Location: Calibration Lab
3. Condition of Item on received: OK
4. Description of Item (UUC):
| Name: | Thermocouple (with Laboratory's reference indicator)
| Sr. No.: | 26072764 | [5 empty] | Make: | YOG
| Type: | K | [5 empty] | Range: | 0 to 800 °C
5. Calibration Standards Used:
Sr. No. | Name |  |  | Unc of Std (±°C) | ID No./ Sr. No. | Certified By |  | Report No. |  |  | Calibration Validity
1 | RTD Sensor with Indicator |  |  | 0.06 | YOG-27, YOG-35 | Nishitronics |  | NI2026/01/0148 |  |  | 2027-01-08
2 | S-Thermocouple with Indicator |  |  | 1.021 | YOG-04, YOG-35 | Nishitronics |  | NI2026/01/0150 |  |  | 2027-01-08
3 | Digital Thermometer |  |  | 0.48 | YOG-36 | Nishitronics |  | NI2026/01/0160 |  |  | 2027-01-07
6. Method Used For Calibration (Work Instruction Ref.): WI – TECH/11
Reference standard No.: EURAMET cg-8
7. Environmental Conditions: |  |  |  | Humidity: 55 ±15 % RH |  |  |  | Ambient Temperature: 25 ±4 °C
8. Calibration Results:
Sr. No. | Standard Value (°C) |  |  | Measured Value on UUC (°C) |  |  | Deviation (°C) |  |  | Expanded Uncertainty (±°C)
1 | 100.08 |  |  | 99.15 |  |  | -0.93 |  |  | 0.61
2 | 400.09 |  |  | 398.58 |  |  | -1.51 |  |  | 0.66
3 | 700.43 |  |  | 698.36 |  |  | -2.07 |  |  | 1.63
UUC: Unit under Calibration
9. Remarks:
1 | A Sticker indicating “CALIBRATION STATUS” has been affixed on the UUC.
2 | The standards used for calibration were calibrated by using reference standard traceable to National / International Standard.
3 | The reported expanded UNCERTAINTY is calculated at 95.45 % of confidence level where coverage factor k=2.
4 | Calibration Points were selected as per customer’s specifications.
5 | The calibration results reported in this certificate are related to this item only and valid at the time of and under stated condition of the measurement.
6 | This certificate shall not be reproduced except in full, without written permission of the laboratory.
7 | Unless otherwise specified in the certificate, the results are applicable only to those products which have been calibrated and do not apply to other products even though declared to be identical.
8 | Above calibration done is meant for scientific & industrial purpose only.
|  |  | Calibrated By : |  |  |  | Approved By : |  |  | Acc. Scope & Cert.
|  |  | Mr. N. H. Bodakhe |  |  |  | Mr. A. A. Kothe
|  |  | (Calibration Engineer) |  |  |  | (Technical Manager)
(----End of the Certificate----)