## -20-0-100
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-9103/1 |  |  |  | Date: |  | 2025-12-21
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9103 |  |  |  | Make: | Fluke
| Sr. No.: | A79198 |  |  |  | Range: | -25 °C to 140 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-08
| Model: | 9103
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9103 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for -20 °C
Sr. No. | Standard Value @ Well-A |  | Standard Value @ Well-B
|  | in °C |  | in °C
1 |  | -20 |  | -19.98
2 |  | -20.01 |  | -19.96
3 |  | -20 |  | -19.95
4 |  | -19.99 |  | -19.96
5 |  | -19.98 |  | -19.94
| Average = | -19.996 | Average = | -19.958
| Stability = | 0.03 | Uniformity = | 0.038 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of 9103 at -20 °C is = |  |  | 0.03 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of 9103 at -20 °C is = |  |  |  | 0.038 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-9103/2 |  |  |  | Date: |  | 2025-12-21
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9103 |  |  |  | Make: | Fluke
| Sr. No.: | A79198 |  |  |  | Range: | -25 °C to 140 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-08
| Model: | 9103
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9103 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 0 °C
Sr. No. | Standard Value @ Well-A |  | Standard Value @ Well-B
|  | in °C |  | in °C
1 |  | 0.01 |  | 0.05
2 |  | 0 |  | 0.04
3 |  | 0.02 |  | 0.05
4 |  | 0.02 |  | 0.04
5 |  | 0.01 |  | 0.05
| Average = | 0.012 | Average = | 0.046
| Stability = | 0.02 | Uniformity = | 0.034 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of 9103 at 0 °C is = |  |  | 0.02 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of 9103 at 0 °C is = |  |  |  | 0.034 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-9103/3 |  |  |  | Date: |  | 2025-12-21
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9103 |  |  |  | Make: | Fluke
| Sr. No.: | A79198 |  |  |  | Range: | -25 °C to 140 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-08
| Model: | 9103
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9103 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 100 °C
Sr. No. | Standard Value @ Well-A |  | Standard Value @ Well-B
|  | in °C |  | in °C
1 |  | 99.99 |  | 100.04
2 |  | 100 |  | 100.03
3 |  | 100.01 |  | 100.05
4 |  | 100.01 |  | 100.03
5 |  | 100 |  | 100.04
| Average = | 100.002 | Average = | 100.038
| Stability = | 0.02 | Uniformity = | 0.036 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of 9103 at 100 °C is = |  |  | 0.02 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of 9103 at 100 °C is = |  |  |  | 0.036 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)

## gradient
GRADIENT REPORT
Report No.: |  | 2512/G-9103/1 |  |  |  | Date: |  | 2025-12-22
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9103 |  |  |  | Make: | Fluke
| Sr. No.: | A79198 |  |  |  | Range: | -25 °C to 140 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-08
| Model: | 9103
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9103 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Temperature Range = Ambient to -20 °C
| Temp. Point |  | Down Cycle Time (Min: Sec) |  | UP Cycle Time (Min: Sec)
| Ambient |  | START |  | 3:41
| 0 |  | 5:17 |  | 1:10
| -20 |  | 17:36 |  | START
Temperature Range = Ambient to 140 °C
| Temp. Point |  | Up Cycle Time (Min: Sec) |  | Down Cycle Time (Min: Sec)
| Ambient |  | START |  | 11:23
| 75 |  | 3:52 |  | 4:50
| 140 |  | 17:51 |  | START
9. Remarks:
Conclusion:
| Temperature gradient of DWB 9103 is |  |  | 17:51 | Min: Sec, which is within acceptable limits
as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)