## 100-250-400
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-9141/1 |  |  |  | Date: |  | 2025-12-23
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9141 |  |  |  | Make: | Fluke
| Sr. No.: | A79661 |  |  |  | Range: | 50 °C to 650 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-09
| Model: | 9141
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9141 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 100 °C
Sr. No. | Standard Value @ Well-A |  | Standard Value @ Well-B
|  | in °C |  | in °C
1 |  | 100.12 |  | 100.3
2 |  | 100.14 |  | 100.28
3 |  | 100.15 |  | 100.32
4 |  | 100.18 |  | 100.33
5 |  | 100.17 |  | 100.34
| Average = | 100.152 | Average = | 100.314
| Stability = | 0.06 | Uniformity = | 0.162 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of 9141 at 100 °C is = |  |  | 0.06 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of 9141 at 100 °C is = |  |  |  | 0.162 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-9141/2 |  |  |  | Date: |  | 2025-12-23
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9141 |  |  |  | Make: | Fluke
| Sr. No.: | A79661 |  |  |  | Range: | 50 °C to 650 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-09
| Model: | 9141
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9141 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 250 °C
Sr. No. | Standard Value @ Well-A |  | Standard Value @ Well-B
|  | in °C |  | in °C
1 |  | 250.28 |  | 250.42
2 |  | 250.25 |  | 250.43
3 |  | 250.21 |  | 250.44
4 |  | 250.24 |  | 250.45
5 |  | 250.23 |  | 250.48
| Average = | 250.242 | Average = | 250.444
| Stability = | 0.07 | Uniformity = | 0.202 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of 9141 at 250 °C is = |  |  | 0.07 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of 9141 at 250 °C is = |  |  |  | 0.202 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-9141/3 |  |  |  | Date: |  | 2025-12-23
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9141 |  |  |  | Make: | Fluke
| Sr. No.: | A79661 |  |  |  | Range: | 50 °C to 650 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-09
| Model: | 9141
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9141 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 400 °C
Sr. No. | Standard Value @ Well-A |  | Standard Value @ Well-B
|  | in °C |  | in °C
1 |  | 400.17 |  | 400.43
2 |  | 400.2 |  | 400.41
3 |  | 400.18 |  | 400.45
4 |  | 400.15 |  | 400.46
5 |  | 400.13 |  | 400.48
| Average = | 400.166 | Average = | 400.446
| Stability = | 0.07 | Uniformity = | 0.28 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of 9141 at 400 °C is = |  |  | 0.07 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of 9141 at 400 °C is = |  |  |  | 0.28 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)

## gradient
GRADIENT REPORT
Report No.: |  | 2512/G-9141/1 |  |  |  | Date: |  | 2025-12-24
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | DWB 9141 |  |  |  | Make: | Fluke
| Sr. No.: | A79661 |  |  |  | Range: | 50 °C to 650 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-09
| Model: | 9141
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of 9141 Dry-well Bath: 4.875 inches (124mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Temperature Range = 50 to 400 °C
| Temp. Point |  | Up Cycle Time (Min: Sec) |  | Down Cycle Time (Min: Sec)
| 50 |  | START |  | 38:34
| 100 |  | 0:32 |  | 21:51
| 200 |  | 1:13 |  | 10:27
| 300 |  | 2:49 |  | 04:28
| 400 |  | 5:43 |  | START
9. Remarks:
Conclusion:
| Temperature gradient of DWB 9141 is |  |  | 05:43 | Min: Sec, which is within acceptable limits
as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)