## 400-800-1200
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-Calsys/1 |  |  |  | Date: |  | 2025-12-17
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Calsys |  |  |  | Make: | Tempsens
| Sr. No.: | 1001 |  |  |  | Range: | 300 °C to 1200 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-10
| Model: | Calsys1200
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | S-Thermocouple with Indicator |  | YOG-04, YOG-35 |  | Nishitronics | NI2025/01/0061 |  | 2026-01-05
5. Immersion Depth of Calsys1200 Bath: 6.299 inches (160mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 400 °C
Sr. No. | Standard Value in °C @ Well-A |  | Standard Value in °C @ Well-B
1 |  | 400.3 |  | 401
2 |  | 400.4 |  | 400.9
3 |  | 400.6 |  | 401
4 |  | 400.5 |  | 401.1
5 |  | 400.7 |  | 401.2
| Average = | 400.5 | Average = | 401.04
| Stability = | 0.4 | Uniformity = | 0.54 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of Calsys1200 at 400 °C is = |  |  | 0.4 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of Calsys at 400 °C is = |  |  |  | 0.54 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-Calsys/2 |  |  |  | Date: |  | 2025-12-18
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Calsys |  |  |  | Make: | Tempsens
| Sr. No.: | 1001 |  |  |  | Range: | 300 °C to 1200 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-10
| Model: | Calsys1200
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | S-Thermocouple with Indicator |  | YOG-04, YOG-35 |  | Nishitronics | NI2025/01/0061 |  | 2026-01-05
5. Immersion Depth of Calsys1200 Bath: 6.299 inches (160mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 800 °C
Sr. No. | Standard Value in °C @ Well-A |  | Standard Value in °C @ Well-B
1 |  | 800.3 |  | 801.1
2 |  | 800.4 |  | 801
3 |  | 800.5 |  | 800.9
4 |  | 800.7 |  | 801.1
5 |  | 800.6 |  | 801.3
| Average = | 800.5 | Average = | 801.08
| Stability = | 0.4 | Uniformity = | 0.58 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of Calsys1200 at 800 °C is = |  |  | 0.4 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of Calsys at 800 °C is = |  |  |  | 0.58 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-Calsys/3 |  |  |  | Date: |  | 2025-12-19
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Calsys |  |  |  | Make: | Tempsens
| Sr. No.: | 1001 |  |  |  | Range: | 300 °C to 1200 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-10
| Model: | Calsys1200
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | S-Thermocouple with Indicator |  | YOG-04, YOG-35 |  | Nishitronics | NI2025/01/0061 |  | 2026-01-05
5. Immersion Depth of Calsys1200 Bath: 6.299 inches (160mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 1200 °C
Sr. No. | Standard Value in °C @ Well-A |  | Standard Value in °C @ Well-B
1 |  | 1200.5 |  | 1201.7
2 |  | 1200.6 |  | 1201.6
3 |  | 1200.9 |  | 1201.4
4 |  | 1201 |  | 1201.2
5 |  | 1200.8 |  | 1201.3
| Average = | 1200.76 | Average = | 1201.44
| Stability = | 0.5 | Uniformity = | 0.68 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of Calsys1200 at 1200°C is = |  |  | 0.5 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of Calsys at 1200°C is = |  |  |  | 0.68 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)

## gradient
GRADIENT REPORT
Report No.: |  | 2512/G-Calsys/1 |  |  |  | Date: |  | 2025-12-21
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Calsys |  |  |  | Make: | Tempsens
| Sr. No.: | 1001 |  |  |  | Range: | 300 °C to 1200 °C
| Type: | Dry Well Calibrator |  |  |  | ID. No.: | YOG-10
| Model: | Calsys1200
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | S-Thermocouple with Indicator |  | YOG-04, YOG-35 |  | Nishitronics | NI2025/01/0061 |  | 2026-01-05
5. Immersion Depth of Calsys1200 Bath: 6.299 inches (160mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Temperature Range = 400 to 1200 °C
| Temp. Point |  | Up Cycle Time (Hr:Min) |  | Down Cycle Time (Hr:Min)
| 400 |  | START |  | 3:43
| 600 |  | 0:17 |  | 1:38
| 800 |  | 0:27 |  | 0:41
| 1000 |  | 0:36 |  | 0:28
| 1200 |  | 0:58 |  | START
9. Remarks:
Conclusion:
| Temperature gradient of Calsys is = |  |  | 00:58 | Hr:Min, which is within acceptable limits
as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)