## 50-100-200
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-OB/1 |  |  |  | Date: |  | 2025-12-24
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Oil Bath |  |  |  | Make: | YOG
| Sr. No.: | OB/01 |  |  |  | Range: | 50 °C to 250 °C
| Type: | Silicon Oil |  |  |  | ID. No.: | YOG-24
| Model: | ---
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of Oil Bath: 4 inches (100mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 50 °C
Sr. No. | Standard Value @ Position-A |  | Standard Value @ Position-B
|  | in °C |  | in °C
1 |  | 49.23 |  | 49.47
2 |  | 49.2 |  | 49.52
3 |  | 49.24 |  | 49.63
4 |  | 49.29 |  | 49.68
5 |  | 49.33 |  | 49.59
| Average = | 49.258 | Average = | 49.578
| Stability = | 0.13 | Uniformity = | 0.32 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of Oil Bath at 50 °C is = |  |  | 0.13 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of Oil Bath at 50 °C is = |  |  |  | 0.32 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-OB/2 |  |  |  | Date: |  | 2025-12-25
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Oil Bath |  |  |  | Make: | YOG
| Sr. No.: | OB/01 |  |  |  | Range: | 50 °C to 250 °C
| Type: | Silicon Oil |  |  |  | ID. No.: | YOG-24
| Model: | ---
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of Oil Bath: 4 inches (100mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 100 °C
Sr. No. | Standard Value @ Position-A |  | Standard Value @ Position-B
|  | in °C |  | in °C
1 |  | 99.07 |  | 99.48
2 |  | 99.01 |  | 99.45
3 |  | 99.08 |  | 99.41
4 |  | 99.16 |  | 99.47
5 |  | 99.14 |  | 99.51
| Average = | 99.092 | Average = | 99.464
| Stability = | 0.15 | Uniformity = | 0.372 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of Oil Bath at 100 °C is = |  |  | 0.15 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of Oil Bath at 100 °C is = |  |  |  | 0.372 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)
STABILITY & UNIFORMITY REPORT
Report No.: |  | 2512/SU-OB/3 |  |  |  | Date: |  | 2025-12-25
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Oil Bath |  |  |  | Make: | YOG
| Sr. No.: | OB/01 |  |  |  | Range: | 50 °C to 250 °C
| Type: | Silicon Oil |  |  |  | ID. No.: | YOG-24
| Model: | ---
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of Oil Bath: 4 inches (100mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Results for 200 °C
Sr. No. | Standard Value @ Position-A |  | Standard Value @ Position-B
|  | in °C |  | in °C
1 |  | 199.06 |  | 199.51
2 |  | 199.1 |  | 199.55
3 |  | 199.03 |  | 199.6
4 |  | 199.11 |  | 199.55
5 |  | 199.19 |  | 199.49
| Average = | 199.098 | Average = | 199.54
| Stability = | 0.16 | Uniformity = | 0.442 | Ref Fig. for Uniformity
9. Remarks:
(Note: All reading of standard have been noted after interval of 5 minute as per WI-TECH/15.)
Conclusion:
1 | Stability of Oil Bath at 200 °C is = |  |  | 0.16 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
2 | Well-to-well Uniformity of Oil Bath at 200 °C is = |  |  |  | 0.442 | °C,
which is within acceptable limits as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)

## gradient
GRADIENT REPORT
Report No.: |  | 2512/G-OB/1 |  |  |  | Date: |  | 2025-12-26
1. Calibration Location: Lab
2. Condition of Item: OK
3. Description of Source:
| Name: | Oil Bath |  |  |  | Make: | YOG
| Sr. No.: | OB/01 |  |  |  | Range: | 50 °C to 250 °C
| Type: | Silicon Oil |  |  |  | ID. No.: | YOG-24
| Model: | ---
4. Calibration Standards Used:
Sr. No. | Name |  | ID No./ Sr. No. |  | Certified By | Report No. |  | Calibration Validity
1 | RTD Sensor with Indicator |  | YOG-27, YOG-35 |  | Nishitronics | NI2025/01/0064 |  | 2026-01-05
5. Immersion Depth of Oil Bath: 4 inches (100mm)
6. Method Used For Checking (Work Instruction Ref.): WI – TECH/15
7. Environmental Conditions:
| Humidity: |  | 55 ±15 % RH |  | Ambient Temperature: |  |  | 25 ±4 °C
8. Results:
Temperature Range = 50 to 200 °C
| Temp. Point |  | Up Cycle Time (Hr:Min) |  | Down Cycle Time (Hr:Min)
| 50 |  | START |  | 2:32
| 100 |  | 0:13 |  | 1:02
| 150 |  | 0:19 |  | 0:29
| 200 |  | 0:44 |  | START
9. Remarks:
Conclusion:
| Temperature gradient of Oil Bath is |  |  | 00:44 | Hr:Min, which is within acceptable limits
as per norms mentioned in WI-TECH/15.
| Checked By : |  |  |  |  | Approved By :
| Mr. N. H. Bodakhe |  |  |  |  | Mr. A. A. Kothe
| (Calibration Engineer) |  |  |  |  | (Technical Manager)
(----End of the Report----)